## Slide 1
Unmixing: It's not magic! © 2025 by David Rach
is licensed under CC BY-SA 4.0
Unmixing: It’s not magic!
By David Rach
Molecular Microbiology and Immunology PhD Candidate
University of Maryland, Baltimore
April 23rd, 2024
Generated by David Rach with Bard, a large language model from Google AI.
Prompt: Cartoon rabbit with rainbow fluorescent spectra.

## Slide 2
A spectral cytometer needs three things
for successful unmixing:
Full-stained Samples
Single-color
Reference Controls
Unstained
Reference Controls

## Slide 3
The Unmixing Screen:

## Slide 4
The Unmixing Screen:
What if I gated like this?
What if I grabbed the entire peak?
What does this even do again!?!?!?
What if I set it on the adjacent detector?

## Slide 5
Rationale:
1) Raw Data
2) ???
3) Unmixed Data
Something something
least squares

## Slide 6
Rationale:
1) Raw Data
2) ???
Something something
least squares

## Slide 7
Goals:
Provide conceptual understanding
Demystify some of underlying math
Help build your own mental model to better troubleshoot issues you may encounter in your own unmixing adventures
Generated by David Rach with Bard, a large language model from Google AI.
Prompt: Computer thinking about how to draw a human.

## Slide 8
A spectral cytometer needs three things
for successful unmixing:
Full-stained Samples
Single-color
Reference Controls
Unstained
Reference Controls

## Slide 9
Let’s go back to the beginning
Unmixing: It's not magic! © 2024 by David Rach
is licensed under CC BY-SA 4.0
Generated by David Rach with Bard, a large language model from Google AI.
Prompt: Lego person jumping into the Back to the Future delorean.

## Slide 10
In our tube, we have a cell

## Slide 11
With some autofluorescent molecules

## Slide 12
Our Aurora shines a laser at it
Excitation

## Slide 13
The cells molecules emit signal back
Excitation
Emission

## Slide 14
Signal is captured across the various detectors
Excitation
Emission

## Slide 15
10000
0
5000
Raw Signal

## Slide 16
Unmixing: It's not magic! © 2024 by David Rach
is licensed under CC BY-SA 4.0
10000
0
5000
Raw Signal

## Slide 17
10000
0
5000
Raw Signal
MFI:4000

## Slide 18
?
What if there were twice as many autofluorescent molecules?

## Slide 19
10000
0
5000
An increase in signal brightness

## Slide 20
10000
0
5000
MFI:6000
An increase in signal brightness

## Slide 21
What if we increased the autofluorescent molecules again?!?
?

## Slide 22
10000
0
5000
An even brighter signal!

## Slide 23
10000
0
5000
MFI:8000
An even brighter signal!

## Slide 24
First essential idea: Brightness
10000
0
5000

## Slide 25
As we record more cells
10000
0
5000

## Slide 26
Accounting for density
10000
0
5000

## Slide 27
And transforming to a logarithmic axis
0
10x4
10x3
10x5
10x6

## Slide 28
Different brightness, but are they the same color?
10000
0
5000
MFI:8000
MFI:6000
MFI:4000

## Slide 29
Normalizing/Scaling by Peak Detector
10000
0
5000
MFI:8000
MFI:5000
MFI:4000

## Slide 30
Divide each detectors measurement by the
peak detectors measurement
10000
0
5000
8000/8000 = 1
5000/8000 = 0.625
4000/8000 = 0.5

## Slide 31
10000
0
5000
8000/8000 = 1
5000/8000 = 0.625
4000/8000 = 0.5
6000/6000 = 1
3000/6000 = 0.5
3750/6000 = 0.625
Divide each detectors measurement by the
peak detectors measurement

## Slide 32
Return the same constants
10000
0
5000
8000/8000 = 1
5000/8000 = 0.625
4000/8000 = 0.5
6000/6000 = 1
4000/4000 = 1
3000/6000 = 0.5
2000/4000 = 0.5
3750/6000 = 0.625
2500/4000 = 0.625

## Slide 33
Second essential idea: Signature
100%
0
50%

## Slide 34
100%
0
50%
Second essential idea: Signature

## Slide 35
10000
0
5000
Similar with Fluorophores (APC)

## Slide 36
100%
0
50%
Similar with Fluorophores (APC)

## Slide 37
Each Fluorophore has its own characteristic signature

## Slide 38
When we gate for Unstained Reference Control
Unstained
Reference Controls
The gate we draw
Selects some cells

## Slide 39
When we gate for Unstained Reference Control
Unstained
Reference Controls
The gate we draw
Selects some cells
The median measurement per detector is taken

## Slide 40
When we gate for Unstained Reference Control
Unstained
Reference Controls
The gate we draw
Selects some cells
The median measurement per detector is taken
And used to generate the reference autofluorescent signature

## Slide 41
A spectral cytometer needs three things
for successful unmixing:
Unstained
Reference Controls

## Slide 42
A spectral cytometer needs three things
for successful unmixing:
Single-color
Reference Controls
Unstained
Reference Controls

## Slide 43
When we gate for Single-color Controls

## Slide 44
When we gate for Single-color Controls
These
gated
cells
Make up both the
positive and
negative peaks

## Slide 45
When we gate for Single-color Controls
These
gated
cells
Make up both the
positive and
negative peaks
The selected detector
Is your x-axis

## Slide 46
When we gate for Single-color Controls
These
gated
cells
Make up both the
positive and
negative peaks
The selected detector
Is your x-axis
The cells within the positive gate
Make up the displayed signature

## Slide 47
And if that wasn’t enough…

## Slide 48
=
+
The signal coming from the positive gate in
our single-color controls is a mixture

## Slide 49
Essential concepts:
Brightness
Signature

## Slide 50
How do we isolate the Fluorophore signature?

## Slide 51
How do we isolate the Fluorophore signature?

## Slide 52
-
= ?
How do we isolate the Fluorophore signature?

## Slide 53
How do we isolate the Fluorophore signature?

## Slide 54
What happens with increased autofluorescence?
?

## Slide 55
What happens if we increase
fluorophore staining?!?

## Slide 56
Very Low

## Slide 57
Low

## Slide 58
Medium

## Slide 59
High

## Slide 60
When autofluorescence is correctly subtracted

## Slide 61
When autofluorescence is correctly subtracted
The Fluorophores signature is the same regardless of brightness

## Slide 62
Difference in cell size:
Same proportion of autofluorescent molecules, but because one cell is larger, more of them overall, thus brighter
Difference in amount of autofluorescence:
Same cell size, just one has more autofluorescent molecules, thus brighter
Variation in autofluorescent molecules:
Same cell size, but express different mixture of autofluorescent molecules, resulting in a different autofluorescent signature
The issue: Variation in autofluorescence

## Slide 63
On dim-staining cells:

## Slide 64
Dim-staining cell: with correct autofluorescence subtraction

## Slide 65
Dim-staining cell: with missed autofluorescence subtraction

## Slide 66
On bright-staining cells:

## Slide 67
Bright-staining cell: with correct autofluorescence subtraction

## Slide 68
Bright-staining cell: with missed autofluorescence subtraction

## Slide 69
Why would there be leftover autofluorescence signal?

## Slide 70
Why would there be leftover autofluorescence signal?
The median measurement from your unstained cells is used.
Example:
UV7 = 2000
V7 = 4000
B3 = 2000

## Slide 71
UV7 = 2000
V7 = 4000
B3 = 2000
Our “average”
autofluorescent cell
for the entire population
Within our single-color reference controls, individual cells may have differences in their expression of autofluorescence molecules.
At the population level, not everything is “average”
UV7 = 1000
V7 = 2000
B3 = 1000
UV7 = 2000
V7 = 4000
B3 = 2000
UV7 = 4000
V7 = 8000
B3 = 4000

## Slide 72
UV7 = 2000
V7 = 4000
B3 = 2000
Leftovers after autofluorescence subtraction
UV7 = 1000
V7 = 2000
B3 = 1000
UV7 = 2000
V7 = 4000
B3 = 2000
UV7 = 4000
V7 = 8000
B3 = 4000
UV7 = 2000
V7 = 4000
B3 = 2000
UV7 = 2000
V7 = 4000
B3 = 2000
UV7 = -1000
V7 = -2000
B3 = -1000
UV7 = 0
V7 = 0
B3 = 0
UV7 = +2000
V7 = +4000
B3 = +2000
=
-
=
-
=
-

## Slide 73
Adventures in Real World Data
(from my recent experiment)

## Slide 74
Substantial leftover autofluorescence affecting signature for
dim marker*fluorophore combinations

## Slide 75
Split the gated cells into 10 groups by percentile, subtracted autofluorescence, and then normalized
Be cautious when using cell controls for dim fluorophores or tertiary markers

## Slide 76
Split the gated cells into 10 groups by percentile, subtracted autofluorescence, and then normalized
Increasing leftover autofluorescence as
you gate further left
Be cautious when using cell controls for dim fluorophores or tertiary markers

## Slide 77
In terms of the fluorophore signature used for unmixing

## Slide 78
In terms of the fluorophore signature used for unmixing
When gating single-color controls, the brightest staining cells are more resilient to signature changes due to leftovers

## Slide 79
One major exception: Antibody aggregates!
Figure from Laura Johnston (formerly University of Chicago) wonderful blogpost! https://voices.uchicago.edu/ucflow/2020/09/30/how-to-identify-bad-flow-cytometry-data-part-1/

## Slide 80
One major exception: Antibody aggregates!
=
-

## Slide 81
Not all problems are due to autofluorescence

## Slide 82
Your fluorophores can also vary in
their signature!!!
Manufacturer Lot Variation
Cross-pollinating BUV and BV fluorophores
Tandem degradation
FRET
And others!
But that’s a problem for Thursday’s Single Color and Autofluorescence Troubleshooting session.

## Slide 83
The single stained control must be as bright or brighter than the fully stained sample.
The exact same fluorophore must be used the stain the control and the fully stained sample.
Background autofluorescence should be the same for the positive and negative control
For now: The three rules for good unmixing controls
https://voices.uchicago.edu/ucflow/flow-cytometry-resources/
Previewing Thursday’s
session

## Slide 84
A spectral cytometer needs three things
for successful unmixing:
Single-color
Reference Controls
Unstained
Reference Controls

## Slide 85
A spectral cytometer needs three things
for successful unmixing:
Single-color
Reference Controls
Unstained
Reference Controls

## Slide 86
Every reference control provides an unique signature
Then what happens?

## Slide 87
10000
0
5000
Each unique signature becomes a reference vector
3
6
3
0
0
3
6
9
6
3
3
6
3
0
0
0
0
0
0
0
0

## Slide 88
5000
3
6
3
0
0
3
6
9
6
3
3
6
3
0
0
0
0
0
0
0
0
10000
0
Each unique signature becomes a reference vector

## Slide 89
3
6
3
0
0
3
6
9
6
3
3
6
3
0
0
0
0
0
0
0
0
Each unique signature becomes a reference vector

## Slide 90
Wait a second…. Vector…

## Slide 91
It’s not magic….

## Slide 92
It’s math!
1) Raw Data
2) Linear Algebra
3) Unmixed Data

## Slide 93
It’s math!
1) Raw Data
2) Linear Algebra
3) Unmixed Data
Which for biologist might as well be magic

## Slide 94
0
0
0
0
0
So lets try to demystify some of underlying math

## Slide 95
A two-detector vector
3
4
1
0
0
2
=
+

## Slide 96
=
+
Finding a linear combination of Blue and Green that gives us the total

## Slide 97
=
+
Finding a linear combination of Blue and Green that gives us the total
3

## Slide 98
=
+
Finding a linear combination of Blue and Green that gives us the total
3
2

## Slide 99
Viewed differently:

## Slide 100
Viewed differently:
3
x

## Slide 101
Viewed differently:
3
x
2
x

## Slide 102
A three-detector vector
=
+
+

## Slide 103
A three-detector vector
=
+
+
3
2
1

## Slide 104
Where David’s Art Skills End
3
4
5

## Slide 105
Lines (or planes) flying through space
3
4
5

## Slide 106
Also explains: Presence or Absence
=
+
+
3
2
1
0
0

## Slide 107
Beyond three-dimensions
=
+
+
+
Each New Fluorophore
Signature

## Slide 108
Final Math Details:
In the above examples, we had the same number of detectors as fluorophores. The matrix generated by collecting the signatures from your single-color controls is a square. This is the case with conventional Flow Cytometry.
In Spectral Flow Cytometry, we have more detectors than Fluorophores. The matrix generated by collecting the signatures from your single-color controls is a rectangle. This is called an overdetermined system. With rare exceptions, these can’t be solved exactly.
However, they can be approximated, which is where some variation of least-square algorithm is used (ordinary least squares in SpectroFlos case)

## Slide 109
If you want to learn more:

## Slide 110
A spectral cytometer needs three things
for successful unmixing:
Single-color
Reference Controls
Unstained
Reference Controls

## Slide 111
Unmixing’s Quandaries
Each cell in your full-stained sample is unmixed individually…
Presence (or absence) of each Fluorophore
Relative expression of each Fluorophore

## Slide 112
Unmixing’s Quandaries
Each cell in your full-stained sample is unmixed individually…
By referencing the unique signatures
from the unstained & single-color reference controls…
….after first subtracting out background autofluorescent signature based on a gated population of unstained cells.
Which are derived from taking the mean/median fluorescence measurement of a gated population of reference cells…
Presence (or absence) of each Fluorophore
Relative expression of each Fluorophore

## Slide 113
Unmixing’s Quandaries
Each cell in your full-stained sample is unmixed individually…
By referencing the unique signatures
from the unstained & single-color reference controls…
….after first subtracting out background autofluorescent signature based on a gated population of unstained cells.
Which are derived from taking the mean/median fluorescence measurement of a gated population of reference cells…
Presence (or absence) of each Fluorophore
Relative expression of each Fluorophore
Variation here will
impact your unmixing results

## Slide 114
The single stained control must be as bright or brighter than the fully stained sample.
The exact same fluorophore must be used the stain the control and the fully stained sample.
Background autofluorescence should be the same for the positive and negative control
For now: The three rules for good unmixing controls
https://voices.uchicago.edu/ucflow/flow-cytometry-resources/

## Slide 115
Questions?